## Slide 1: Setup of laser tracker on DX side
NOTE:
X axis along M1s centers, towards SX
Z axis along M1-M2 centers, towards M2

## Slide 2: Taking right front bend gregorian rotator as reference
Corner cube and nest glued to rotator flange.
Rotator moved and circle fitted to identifiy the axis.

## Slide 3: Nest and optic installed on dichroic mount

## Slide 4: Measurements in the parked and working position
Final error on dichroic tilt: -0,0046deg (wrt 15deg)
Final error on dichroic tip: 0,0051deg
Dichroic center offset wrt rotator axis: [0,0,-1]mm

## Slide 5: 3 M6 holes position measured on DX table

## Slide 6: Position of the M6 holes on DX table (to be compared with Matthias model)
P1: [-1077.2, -1043.6]
P2: [-1217.5, -1618.0]
P3: [-780.5, -1079.0]
Reference: intersection point between the rotator axis and the rotator flange
Table plane after repositioning to the end of the slot is 147mm below rotator axis (nominal 150mm)

## Slide 7
Position of the M6 holes on SX table (to be compared with Matthias model)
P1: [1092.9, -1059.8]
P2: [1232.4, -1633.3]
P3: [794.8, -1095.3]
Reference: intersection point between the rotator axis and the rotator flange
Table plane is 153mm below rotator axis (nominal 150mm)
-> To be repositioned: actually is at the end of the slot.

## Slide 8: Rotator axis visualization on DX table: flat mirror installed on front of AGW

## Slide 9: Rotator axis visualization on DX table: flat mirror aligned perpedicular to rotator axis
0deg rotator angle
180deg rotator angle
Laser beam projected across M1 to flat mirror and reflected on the dome

## Slide 10: Rotator axis visualization on DX table: rotator axis visualized on the flat mirror
Laser beam projected across M1 and position marked at 0, 90, 180, 270deg of rotator angle
Center of 4 spots is the rotator axis (center of the cross)

## Slide 11: Rotator axis visualization on DX table: laser beam projected from table and autocollimated through dichroic
LGSW table
laser
1’’ mirror
dichroic
Flat mirror on AGW

## Slide 12: Rotator axis visualization on DX table: laser beam reflected by dichroic and fold mirror on DX table
LGSW table
laser
1’’ mirror
dichroic
Fold mirror

## Slide 13: References installed for LGSW positioning
To be done: set beam height to 150mm tilting dichroic and fold mirror